Commissioning Document for the TITAN Penning Trap Mass Spectrometer

Vladimir Ryjkov

10th August 2007

      - Commission 4T Superconducting Solenoid Magnet.

- Install the magnet, verify cryostat integrity, cool down

- Charge the magnet

- Conduct the test quench of the magnet

- Verify operation of the shimming coils

- Conduct NMR probe mapping and shimming of the magnet

- Conduct radial Hall probe mapping to determine location of the central field line

- Test field decay in the persistent mode

- Check cryogen consumption

Magnet Commissioned by:___________________________________

Signed: _______________________	Date:_______________________

      - Commission Vacuum System

        - Assemble ion trap and optics, place them into the titanium vacuum chamber

        - Place titanium vacuum chamber into the bore of the superconducting solenoid

        - Assemble the rest of the vacuum vessels, install vacuum pumps, gauges and valves.

        - Leak check the vacuum vessels and joints

        - Install Vacuum Control System

Vacuum System Commissioned by: :_________________________________

Signed: _______________________	Date:_______________________

      - Commission Trap Optics Supplies and Controls

- Install optics supplies in the racks

- Install optics control system

- Turn on the supplies and check operation of the controls and interlocks

Trap Optics Controls System

commissioned  by: _______________________________________

Signed: _______________________	Date:_______________________

      - Install and Commission MCP Detector

- Install and wire the MCP detector

- Install detector electronics

- Bias the detector and verify detector operation

MCP detector installed and tested by: ___________________________________

Signed: _______________________	Date:_______________________

      - MIDAS DAQ System Installation

- Install the DAQ system and electronics into the racks

- Install MIDAS DAQ system and frontend software and verify their operation

- Verify the programmable pulse generator operation

- Verify TDC operation

- Verify AWG operation

- Verify operation of the MSCB RF generator controls

MIDAS DAQ system checked by: ___________________________________

Signed: _______________________	Date:_______________________

      - High Voltage Pulser Installation

- Install high voltage pulser

- Verify its operation

- Connect the high voltage pulser to the pulsed drift tube electrode

- Verify the operation of the pulser with connected drift tube

- Check and record transient voltages on the adjacent electrodes generated by pulsing of the drift tube

High voltage pulser tested by: ___________________________________

Signed: _______________________	Date:_______________________

      - Electron Source Installation

- Assemble and install the electron source

- Activate the cathode and record the maximum current

Electron source installed by: ___________________________________

Signed: _______________________	Date:_______________________

      - Penning Trap Tests with Electron Source

- Extract the beam from the electron source.

- Detect on the MCP the ions generated by the electron beam in the vacuum chamber

- Pulse the electron beam by controlling the voltage on the electron source aperture plate and observe the ion signal pulses on the MCP

- Verify the ion generation and capture in the Penning trap by pulsing the electron beam and, after a delay, pulsing the trap electrodes open and observing the ion signal on the MCP

- Set up DAQ sequences for the electron source tests

- Conduct time-of-flight (TOF), magnetron excitation, axial excitation, cyclotron excitation tests on the ions captured in the Penning trap.

Penning trap electron source tests

conducted by: ___________________________________

Signed: _______________________	Date:_______________________

      - RF amplifier system installation

- Install RF amplifiers, switching and diagnostic components, and cabling

- Verify the operation of the RF amplifier system

- At different power levels, check and record the phase and amplitude characteristics of different amplifier channels

- Calibrate the RF detector circuits

- Connect the RF system to the split guard electrode feedthrus and verify system operation

RF amplifier system installed

and checked by: ___________________________________

Signed: _______________________	Date:_______________________

      - RFQ buncher capture tests

- Direct the RFQ ion bunches into the magnetic field and detect them on the other side of the magnet with the MCP detector

- Set up the DAQ sequence for the RFQ capture test

- Triggering the RFQ extraction from the MIDAS DAQ system, pulse the MPET drift tube. Adjusting the MPET drift tube pulse down delay, observe change in time of arrival of ion bunches onto the MCP detector.

- Tune the MPET drift tube delay for correct energy tuning of the RFQ ion bunches

- Adjust the MPET drift tube voltage so that the ion bunch energy is suitable for the capture in the trap (under 10eV).

- Pulse the trap electrodes to capture the ions, and, after a delay, to release them towards the MCP detector.

- Verify the ion capture by observing on the MCP the signal generated by the ions captured and then extracted from the trap

Capture of RFQ bunches verified by: ___________________________________

Signed: _______________________	Date:_______________________

      - Mass measurements on the ions from RFQ buncher

- Set up the DAQ sequences for the RF excitation tests

- Triggering the RFQ extraction, verify ion capture

- Conduct time-of-flight, axial excitation, magnetron and cyclotron excitation tests on the captured ions

- Using the MPET beamline optics and the Lorentz steerer, optimize the ion injection into the trap

- Conduct mass measurements on the ions from the RFQ buncher

Mass measurements set up

and conducted by: ___________________________________

Signed: _______________________	Date:_______________________

TITAN Penning trap mass spectrometer

commissioned  by: ___________________________________

Signed: _______________________	Date:_______________________